Assessment of inter-rater and intra-rater reliability of cervical range of motion (CROM) goniometer

ABSTRACT

Purpose: The current study was designed to assess inter-rater and intra-rater validity of cervical range of motion measurements performed with CROM goniometer.

Material and methods: The study involved 95 healthy university students (31 males and 64 females) aged 20-24 years. Two examiners performed measurements of cervical range of motion using CROM goniometer. The same subjects were examined again after two weeks, in the same conditions.  The results acquired by one rater during the first and the second examination were compared for reproducibility, while the results obtained by the two examiners were compared to assess validity and reliability of the tool.
Cronbach’s alpha was applied to determine intra-rater reliability, and the values of correlations were used to assess the inter-rater agreement.

Results: Analysis of the results showed both intra-rater and inter-rater agreement in all the measures of cervical range of motion. The highest intra-rater and inter-rater concordance was observed in the measure of extension. Intra-rater agreement for Examiner 1 was reflected by Cronbach’s alpha = 0.85, and for Examiner 2 by Cronbach’s alpha = 0.89. As for the inter-rater agreement in the measure of extension, the value of correlation in both the first and the second measurement amounted to r= 0.89.

Conclusions: Measurements performed with CROM goniometer show inter-rater and intra-rater agreement in assessments of cervical range of motion. CROM goniometer can be recommended for use in daily clinical practice.

Keywords: cervical spine, range of motion, CROM goniometer

Introduction

The most common dysfunctions of the locomotor system include dysfunction of the cervical spine [1,2], which is recognised as the most mobile segment of the spine [3]. Both the biomechanics and anatomy of the cervical spine contribute to the risk of strain, particularly affecting the intervertebral disc. This leads to various ailments generally referred to as neck pain syndrome [4]. Cervical spine disorders are common and they significantly impact quality of life, also constituting a problem for healthcare systems. Patients with cervical spine disorders usually experience pain, and reduced active and passive cervical ranges of motion. Recovery of physiological mobility is of key importance in cervical spine disorders [5].

The range of motion is a commonly accepted parameter used in assessment and evaluation of spine movements [6], yet it is also thought to be extremely difficult to access accurately due to the complex anatomy and the associated movements. The cervical range of motion is frequently applied as a baseline and outcome measure in documenting the effects of interventions, and in adjusting treatment plans in physiotherapy [7].  Physiotherapists and physicians can routinely perform assessment of active cervical range of motion in order to estimate the level of mobility limitation and its association with pain, and to evaluate effects of treatment [8]. Many tools have been developed to enable measurement of cervical mobility, ranging from simple visual examination to the complex three-dimensional mobility assessment [5].

To be clinically useful, a measure of cervical range of motion must be confirmed for validity, reliability and reproducibility [9]. Reproducibility reflects to what degree the results identified with a given device are stable over time, if the relevant subjects do not change. Reliability, the most frequently evaluated aspect, is a quality showing to what extent patients are differentiated, despite measurement errors [8]. Many other tools have been described in the literature as possible options for clinical evaluation when choosing an adequate method of assessing the cervical range of motion, yet there is no agreement as to which is the gold standard method for measurement of range of motion [10, 11].

Cervical Range of Motion (CROM) goniometer is a clinically available tool enabling straightforward measurement of neck mobility [7]. Review of literature shows there are a lot of studies focusing on CROM goniometer [5,7,8,10].  However, in Poland so far there are no data showing validity of CROM goniometer. This observation provided a motivation for the present study, and this article is the first research report related to this issue in our country.

The purpose of the study was to assess intra-rater and inter-rater validity of cervical range of motion measurements performed with CROM goniometer.

Material and methods

Participants

The study involved 95 healthy subjects, students of physiotherapy, aged 20-24 years, with no cervical spine dysfunctions, which was confirmed by medical examination. The subjects’ mean age was 21 years. The group included 31 males and 64 females. Patients with pain symptoms, spinal disorders and concurrent spinal mobility illnesses were excluded from the research.

Measurements

The cervical range of motion was examined using CROM goniometer [7, 12, 13]. The examinations were performed in a Spine Kinesiology facility. The range of motion measurement was carried out with CROM goniometer by two independent examiners, during the same time and in the same conditions. To ensure the greatest reliability of the measurements the examiners did not contact each other. The examiners are referred to as 1A and 2P.

Examination of mobility was repeated two weeks later in the same subjects and in the same conditions; the examination was carried out by the same examiners to enable evaluation of reproducibility of the measurements performed with CROM goniometer (after two weeks: examiner 3A, and examiner 4P).

The measurements of the cervical range of motion, carried out with CROM goniometer, were performed in all the planes of motion, with an accuracy of 1° (Figure 1). The goniometer was placed on the subject’s head and a magnetic collar was attached to their shoulders; it was always arranged in the same position with respect to the magnetic pole.  During all the measurements the subjects were seated, with their back straight, and looking ahead; their feet were supported on the floor. The subjects remained in the same sitting position from the moment the recording procedure was started. The measurement was performed during “half a cycle” of the movement, i.e. from “the neutral” position to the final range of motion in a given plane. Each measurement of the active range of motion was preceded with instructions, whereby the examiner presented the movement to be performed by the subject. Subsequently the subject performed a practice trial. During each trial the movement was executed to reflect the subject’s maximum capabilities, i.e. up to the moment they could feel pain or stretching of the tissues, or until substitution of movement occurred. When the maximum movement was performed in one direction, the final position was identified and recorded [7].

Figure 1. Measurement of the range of cervical spine motion using the CROM goniometer.

Statistical analysis

Intra-rater reliability was assessed by comparing two measurements carried out by one examiner. Assessment of intra-rater reliability was performed using Cronbach’s alpha coefficient whose value should not be lower than 0.6.

Evaluation of inter-rater agreement involved comparison of measurements performed by two examiners, during the first session (1A versus 2P), and during the second session two weeks later (3A vs 4P). Assessment of inter-rater agreement was based on the value of correlation (r), which should not be lower than 0.5.

The sample selection calculator was used to calculate the minimum sample size, taking into account the number of people studying physiotherapy. A fraction size of 0.8 was applied, with a maximum error of 6%, and as a result a sample size of 92 people was determined.

Results

The flow of the subjects through the study is shown in Figure 2. No adverse events were observed during the study.

Based on the results of the above examinations, Table 1 presents the findings related to intra-rater and inter-rater reliability in the measure of flexion. It was shown that intra-rater reliability in the measure of flexion for Examiner A was reflected by Cronbach’s alpha = 0.75, and in the case of Examiner P was expressed by Cronbach’s alpha = 0.64. Inter-rater agreement in the measure of flexion during Exam I performed by Examiners 1A and 2P is shown by the value r=0.72 and in the case of Exam II, performed by Examiners 3A and 4P is indicated by the value r=0.5.

Analysis of the results related to intra-rater and inter-rater reliability in the measure of extension showed that intra-rater reliability in the measure of extension for Examiner A was reflected by Cronbach’s alpha = 0.85, and in the case of Examiner P was expressed by Cronbach’s alpha =0.89. Inter-rater agreement in the measure of extension during Exam I performed by Examiners 1A and 2P and in Exam II, performed by Examiners 3A and 4P is shown by the value r=0.89. Results related to intra-rater and inter-rater reliability in the measure of extension are presented in Table 2.

Table 3 presents results related to intra-rater and inter-rater reliability in the measure of left lateral flexion. It was shown that intra-rater reliability in the measure of left lateral flexion for Examiner A was reflected by Cronbach’s alpha = 0.7, and in the case of Examiner P was expressed by Cronbach’s alpha =0.75. Inter-rater agreement in the measure of left lateral flexion during Exam I performed by Examiners 1A and 2P is shown by the value r=0.61 and in the case of Exam II, performed by Examiners 3A and 4P, is indicated by the value r=0.71.

Analysis of the results related to intra-rater reliability and inter-rater agreement in the measure of right lateral flexion showed that intra-rater reliability in the measure of right lateral flexion for Examiner A was reflected by Cronbach’s alpha = 0.72, and in the case of Examiner P was expressed by Cronbach’s alpha =0.64. Inter-rater agreement in the measure of right lateral flexion during Exam I performed by Examiners 1A and 2P is shown by the value r=0.82 and in the case of Exam II, performed by Examiners 3A and 4P, is indicated by the value r=0.72 (Table 4).

Analysis of the results related to intra-rater reliability and inter-rater agreement in the measure of left rotation showed that intra-rater reliability in the measure of left rotation for Examiner A was reflected by Cronbach’s alpha = 0.65, and in the case of Examiner P was expressed by Cronbach’s alpha =0.66. Inter-rater agreement in the measure of left rotation during Exam I performed by Examiners 1A and 2P is shown by the value r=0.67 and in the case of Exam II, performed by Examiners 3A and 4P, is indicated by the value r=0.74 (Table 5).

Comparison of the results related to intra-rater and inter-rater reliability in the measure of right rotation showed that intra-rater reliability in the measure of right rotation for Examiner A was reflected by Cronbach’s alpha = 0.67, and in the case of Examiner P was expressed by Cronbach’s alpha =0.72. Inter-rater agreement in the measure of right rotation during Exam I performed by Examiners 1A and 2P is shown by the value r=0.66 and in the case of Exam II, performed by Examiners 3A and 4P, is indicated by the value r=0.72 (Table 6).

# Discussion

The study was carried out to assess reliability and reproducibility of CROM goniometer measurements. The assessments were performed in three planes, in subjects showing no cervical spine dysfunctions. The current findings show good inter-rater and intra-rater reliability of CROM goniometer. Cronbach’s alpha coefficient in all the measures exceeded the required value, reflecting good intra-rater reliability of CROM goniometer. Likewise, the value of correlations in all the measurements exceeded the minimum necessary to confirm inter-rater agreement of CROM goniometer measurements.

There are numerous reports in the literature discussing reliability of CROM [5,7,8,10]. In Poland, however, there are no publications related to this subject matter. This observation provided a motivation to carry out the present study on polish ground.

Chantal et al. assessed reliability, validity and responsiveness of devices applied in measurement of cervical mobility, including CROM goniometer. Their research shows that, given its well-established clinimetric properties, the device may easily be used in daily clinical practice [8]. According to Audette et al., given the emphasis to the importance of objective tools, CROM goniometer may be recommended for use in assessing cervical mobility both for the needs of research and in rehabilitation contexts. The reported findings confirmed good reliability of CROM goniometer; the authors also assessed error of measurements and reported good test-retest reliability of the device [7]. Fletcher and Bandy published a study focusing on intra-rater reliability of the tool in subjects with and without neck pain. Their findings are consistent with those reported in the present study. They documented sufficient reproducibility of CROM goniometer measurements, demonstrating that the device can effectively be used in daily clinical practice in patients with neck pain and in subjects with no symptoms of cervical problems [10]. Mangone et al. presented comparable results related to reliability and reproducibility of cervical range of motion measurements performed with CROM goniometer. The study involved healthy individuals, with no spinal dysfunctions. These authors emphasise importance of such factors as the affordability of a CROM device and the simplicity of its operation, as well as its portability owing to the small size of the tool [5]. Rondoni et al. published a review which was designed to assess intra-rater and inter-rater reliability of active cervical range of motion measurements performed with expensive technological equipment as compared to affordable common use devices. The study involved subjects with nonspecific neck pain. The authors additionally investigated whether the reliability of the measuring device was influenced by the plane of the assessed movement. Conclusions from this review suggest that the reliability of the assessment is not significantly affected by the type of the measuring device, hence a comparison of more expensive and more affordable devices showed no significant effects. Likewise, the plane of movement does not influence the results in cervical range of motion measurements [13].

The current findings show satisfactory results in the movements for all the planes, i.e. sagittal, frontal and transverse, while the best results were observed in the assessment of extension, in terms of both inter-rater and intra-rater agreement. Likewise, highly significant statistical correspondence was found in the measurement of left lateral flexion. The lowest values could be observed in measurements of left and right rotation, however, even in this case the results were higher than the minimum values required for inter-rater and intra-rater reliability.

Similarly, Kenneth et al. demonstrated that the most consistent CROM measurements are found in assessments of flexion, extension and lateral flexion performed with inclinometer. The measurements were carried out by 33 practising physiotherapists who examined three persons. The measures of the range of flexion, extension and lateral flexion were compared with results shown in X-ray pictures, while the ranges of rotation were compared with axial tomography scans [14]. In another study, Kenneth et al. also obtained good results for flexion and extension, in assessment of inter-rater and intra-rater reliability of upper cervical motion measurements performed with CROM goniometer in a group of 30 healthy individuals aged from 23 to 37 years [15]. Likewise, Pucci et al. in a study involving 20 healthy subjects and assessing inter-rater and intra-rater reliability of CROM goniometer observed the best results in the measures of flexion and extension [16].

While investigating other methods of cervical range of motion measurement, researchers point out that clinical usefulness of such complex devices as electrogoniometers and digital inclinometers is doubtful due to practical reasons and costs involved. Dynamic X-ray examination of intervertebral motion for years was recognised as the “gold standard” in assessment of cervical range of motion, today however applicability of this invasive method is limited due to the small number of X-ray images which can be obtained while ensuring minimum exposure of the patient to radiation. Another method, i.e. three-dimensional kinematic analysis is also known for the accuracy of the calculations, however it is also expensive and it can be used only in laboratory settings, and require experienced personnel [17,5,3].

The study limitation is to narrow down the researched patients to one age group. Follow up research encompassing various age groups, especially 50 plus one seems to be necessary to be conducted.

In view of the above, it appears from the analysis of the acquired findings that assessment of cervical motion with CROM goniometer results in highly reliable and reproducible measurements which are easy to perform and do not require significant financial resources. Notably, this is the first scientific report related to this subject matter in Poland and it provides clinicians with an additional tool which can be applied in the daily practice. In order to validate the findings, further studies should be conducted in groups of subjects with various cervical spine dysfunctions.

# Conclusion

Measurements performed with CROM goniometer show inter-rater and intra-rater agreement in assessments of cervical range of motion.

CROM goniometer can be recognised as an alternative tool which can be recommended for use in daily clinical practice.

# REFERENCES

- Wolański W, Tejszerska D (2009) Biomechanical analysis of the cervical spine 
in a situation of stabilization; Engineering Modeling; 7(38): 295-300

- Chaves TC, Nagamine HM, Belli JFC et al (2008) Reliability of fleximetry
 and goniometry for assessing cervical range of motion among children: Brazilian Journal of Physical Therapy; 12(4):283-289

- Porzych P, Ratuszek-Sadowska D, Pyskir M et al  (2016) Cervical vertebral column - mobility and selected ways of measurement - a literature review; Journal of Education; 6(6):505-516

- Mikołajczyk E, Jankowicz-Szymańska A, Guzy G et al (2013)The impact of comprehensive physiotherapy on the functional state of patients with pain in the cervical spine; Hygeia Public Health; 48(1):73-79

- Mangone M,  Bernetti A, Germanotta M et al (2018) Reliability of the cervical spine device for the assessment of cervical spine range of motion in asymptomatic participants; J Manipulative Physiol Ther; 41:342-349

- Jonas R,  Demmelmaier R, Hacker SP et al (2018) Comparison of three-dimensional helical axes of the cervical spine between in vitro and in vivo testing; Spine Journal; 18(3):515–524

- Audette I, Dumas JP, Côté JN et al (2010) Validity and between-day reliability
 of the cervical range of motion (crom) device; J Orthop Sports Phys Ther ; 40(5):318-23

- Chantal HP, de Koning, Sylvia P, van den Heuvel, J Bart Staal  et al (2008) Clinimetric evaluation of active range of motion measures in patients with non-specific neck pain: a systematic review; Eur Spine J; 17:905–921

- Williams M A, McCarthy CJ, Chorti A et al (2010)A systematic review of reliability and validity studies of methods for measuring active and passive cervical range of motion; J Manipulative Physiol Ther; 33(2):138-55

- Fletcher JP, Bandy WD (2008) Intrarater reliability of CROM measurement of cervical spine active range of motion in persons with and without neck pain; J Orthop Sports Phys Ther; 38(10):640-5

- Lachtman DS, Bartha DA, Beltran MM et al (2015) Rater reliability and concurrent validity of single and dual bubble inclinometry to assess cervical lateral flexion; J Manipulative Physiol Ther; 38(8):572-80

- Davies M, Moore IS, Moran P et al (2016) Cervical range of motion, cervical
 and shoulder strength in senior versus age-grade Rugby Union International front-row forwards; Phys Ther Sport.; 19:36-42

- Rondoni A,  Rossettini G,  Ristori D et al (2017) Intrarater and inter-rater reliability of active cervical range of motion in patients with nonspecific neck pain measured with technological and common use devices: a systematic review with meta-regression; J Manipulative Physiol Ther.; 40(8):597-608

- Bush K.W, Collins N, Portman L, Tillett N (2000) Validity and Intertester Reliability of Cervical Range of Motion Using Inclinometer Measurements; The Journal of manual & manipulative therapy 8(2):52-61.

- Suzette K.D, Brodeur M, Croteau S R ( 2013) Reliability of the Cervical Range of Motion Device in Measuring Upper Cervical Motion.; The Journal of manual & manipulative therapy 6(1).

- Capuano0Pucci D, Rheault W, Auka J, Bracke M, Day R, Pastrick M (1991)Intratester and intertester reliability of the cervical range of motion device. Arch Phys Med Rehab; 72:338–340.]

- Carvalho GF, Chaves TC, Gonçalves MC et al (2014) Comparison between neck pain disability and cervical range of motion in patients with episodic and chronic migraine: a cross-sectional study; J Manipulative Physiol Ther; 37(9):641-6

Abbreviations

CROM: cervical range of motion

Declarations

Consent for publication

Not applicable

Competing interests

The authors declare that they have no competing interests.

Study protocol

The full study protocol can be obtained from the corresponding author

Funding

This research did not receive any specific grant from funding agencies in the public, commercial, or not-for-profit sectors.

Acknowledgements

Not applicable